## Slide 1
First Grade- 2024-2025
The Basic of Fluoroscopy
Instructor: Prof. Dr. Ronak T. Ali
Tishk International University
Faculty of Applied Science
Medical Technical Radiology
6

## Slide 3
Study of moving body structures.
  Similar to an x-ray
Fluoroscopy is also an imaging tool.
  Allows physicians to look at various body systems.

## Slide 4
Show continuous x- ray image
Plays out like a movie
  Images taken quickly allow for this to happen
Shows movement of body parts
  Also shows
instruments or dye

## Slide 5
The primary function of a fluoroscope is to perform dynamic studies; that is, the fluoroscope is used to visualize the motion of internal structures and fluids.
The purpose of this technique is to get real-time and moving images of the insides of a person by way of the fluoroscope.
If something is observed that the radiologist would like to preserve for later study, a radiograph can be made with little interruption of fluoroscopic examination. Such a radiograph is known as spot film.

## Slide 6: Scattered radiation
There are three types of X –rays:
primary X –rays : are released from the X -ray tube
scattered X –rays : are scattered from the primary substance and occur as a result of collision of electrons;
residual radiation : is X -rays that pass through the patient and strengthen the image.

## Slide 7: Scattered radiation
If the X -ray tube is on the patient and at eye level to the physician performing the scanning, the physician ’s face is relatively close to the X -ray tube and is unprotected against radiation leaks scattered from the tube.
The image booster, if it is positioned above the patient, also serves as a barrier to protect the physician ’s face.
in this position the tube is placed below the patient, back scattering also occurs toward the ground

## Slide 8: Scattered radiation
Fluoroscopy devices that automatically determine the beam settings also automatically adjust the tube potential (kV) with increasing thickness of the patient ’s body
With an increase in tube potential, the force of the primary beam increases and scattering radiation will be more penetrating

## Slide 9: Fluoroscopy usage types
Continuous fluoroscopy:
  found in all the old machines, and is still a format that is used.
  Fluoroscopy continues as long as the pedal is held down.
  This format is used incorrectly by many physicians.
  From the aspect of radiation safety, this is very dangerous.
High-speed fluoroscopy:
  used in situations where normal fluoroscopy is inadequate.
  it is not recommended.
Pulsed fluoroscopy:
  The image speed and frames per second rate (fps) is selected.
  This is the most advantageous format.

## Slide 10
Takes a continuing stream of x-ray images
  Approximately 25- 30 images per second
  Images are viewed on a monitor
    Sort of like a
television screen.

## Slide 11
Amount of radiation needed various
  Based on procedure
Important characteristic of Fluoroscopy
  Sensitivity
    Amount of exposure needed to create an image
Non-intensified Fluoroscopy
  Uses a fluorescent screen only for a receptor
    Should not be used because of excessive exposure

## Slide 12: Components of Fluoroscope
x-ray generator
x-ray tube
collimator
filters
patient table
grid
image intensifier
optical coupling
television system
image recording

## Slide 13
Used in a variety of procedures
  Examples include:
    Orthopedic Surgery
      Observe fractures and healing bones
    Catheter Insertion
      Direct catheter placement (Angiography/Angioplasty)
    Barium X-Rays
      Observe movement through GI tract
    Blood Flow Studies
      View blood flow to organs

## Slide 14
Injections into the knees
  Viscosupplementation injections
Locating foreign bodies
Percutaneous Vertebroplasty
  Treating compressed fractures of the spine
Injections into joints or spine
  Image-guided anesthetic injections

## Slide 15
Fluoroscopy used alone
  Gives physician opportunity to see movement in the intestines
    Barium moves through them during procedure

## Slide 16
Fluoroscopy used alone
  Aids physicians in inserting a catheter
  Also aids them in detecting blockages in arteries
  Physicians can see the flow of blood

## Slide 17
Because Fluoroscopy is an x-ray machine, it has the same risks as other x-ray machines.
Two major risks
  There is a small possibility of developing cancer due to the exposure to the radiation
  Injuries such as burns caused by the radiation
Benefit
  If a patient is in need of a Fluoroscopy, the benefit outweighs the minute risks

## Slide 18
Insertion of an IV into patient’s hand or arm
Patient moved onto x-ray table
Additional line may be inserted for catheter procedures
X-Ray scanner used to create fluoroscopic images of the body
Dye may be injected into the IV at this point
Type of care will be decided on after the procedure has finished

## Slide 19
Continuous x-ray passes through the body
Beam passes onto a television monitor
Body part and motion can be seen in great detail

## Slide 20
Two main things to consider
  Area most exposed
  Total radiation absorbed
Area Most Exposed
  Highest absorbed dose
    In the general area, as well as specific organs
Total Radiation absorbed
  Can result in injuries
    Burns, etc.
    Caused by prolong exposure

## Slide 21
Fluoroscopy is also an imaging tool.
  Allows physicians to look at various body systems.
Shows movement of body parts.
  Also shows instruments or dye
Takes a continuing stream of x-ray images.
  Approximately 25-30 images per second
Used in a variety of procedures
  Orthopedic Surgery
  Catheter Insertion
  Barium X-Rays
  Blood Flow Studies

## Slide 22
Two major risks
  There is a small possibility of developing cancer due to the exposure to the radiation
  Injuries such as burns caused by the radiation
Benefit outweighs the risks
Precise procedure
  Plenty of steps followed to ensure a successful procedure
Plenty to consider during procedure
  Area Most Exposed
  Total Radiation absorbed

## Slide 23: Effects of radiation on biological systems
Radiation effects on biological systems are divided into two Groups:
Determining (direct, definitive) effect
Cytocastic
(indirect, nondefinitive) effect

## Slide 24: Cytocastic (indirect, nondefinitive) effect
The International Committee on Radiation Safety Protection ICRP has produced estimates of the maximum permissible dose (MPD) of annual radiation to various organs
Exposure below these levels is unlikely to
lead to any significant effects
the ICRP recommends that workers should not receive more than 10% of the MPD

## Slide 25: Radiation exposure and dose units

## Slide 26: Yearly maximum radiation dosages

## Slide 27: Radiation safety
necessity for imaging or the procedure should be
assessed:
  effectiveness of the procedure
  personal dose limits
  risk assessment
radiation protection should take into account three basic rules:
  Time
  Distance
  use of protective materials (shielding)

## Slide 28: Fundamental principles of radiation protection
Hold the tube low.
Keep away from sources of radiation.
Consider the scattering profile.
Keep the beam time brief.
Set the appropriate geometry
Eliminate X -rays from the outside
The device should undergo technical checks
Wear protective clothing.

## Slide 29
During routine use in the anterior-posterior plane, the x-ray tube (source) should be positioned below the patient and the detector above the patient to minimize radiation exposure to both the patient and the practitioner.

## Slide 30
The oblique projection results in markedly increased exposure to the practitioner

## Slide 31
During use in the lateral projection, the practitioner should step completely behind the x-ray tube (source) to minimize radiation
exposure.
When it is necessary to work close to the patient during lateral fluoroscopy, the practitioner should step away from the x-ray tube and move to the side of the table opposite the x- ray tube to minimize exposure
DR MOHSEN ABAD 55

## Slide 32: Radiation exposure to both the patient and the practitioner is dramatically increased when the x-ray tube (source) is inverted above the patient.
Some practitioners invert the (-arm to allow for more extreme lateral angle (e.g., rotation beyond 35 to 45 degrees oblique to the side opposite the (-arm is not possible without inverting the (-arm on some units). Radiation exposure can be reduced by rotating the patient on the table and keeping the x-ray source below the table.

## Slide 33: Keep away from sources of radiation
what is a safe distance?
stand 3 meters away from the table to reduce the beam level that can be absorbed to an acceptable level.
best place to stand for staff is where the X -ray scatters twice before it reaches them
radiation intensity decreases approximately 1000 times with each scattering event
in other words one -millionth of the original value

## Slide 34: Every fluoroscopy operator should do the following
Track the beam time:
– The C -arms of fluoroscopy devices have an irradiation alarm that sounds for every 5 minutes of irradiation
Save the last image:
  Storing the last image in memory in this manner will significantly reduce the radiation dose
  Pressing the pedal for a long time is not a factor in increasing quality the image brightness, or contrast.
Use the pulse mode:
  two different modes, continuous and pulse.
  In pulse –mode imaging, intermittent versus continuous radiation.
  This mode reduces the irradiation time by up to 2 to 4
times.

## Slide 35: Every fluoroscopy operator should do the following
Limit the beam size (collimation)
  If possible, should restrict the irradiated area by reducing the diaphragm or with the collimator.
  This process reduces the amount of scattered radiation and
improves the image quality
Set the appropriate geometry:
  If the X -ray tube is brought too close to the patient, it can
cause skin burns during long -term applications
  if the X-ray tube is too far from the patient, this can increase the size of the image
  As a rule, doubling the focus distance increases the rate of
exposure to radiation fourfold.

## Slide 36: Every fluoroscopy operator should do the following
Eliminate X -rays from the outside:
  Outside light barriers interacting with the fluoroscopy device can block the image details from being seen.
  In such a case, the fluoroscopic image needs to be enlarged or
the amount of scattered light increased.
The device should undergo technical checks:
  Technical control is very important
  image intensifier ages
  automatic brightness control mode

## Slide 37: Every fluoroscopy operator should do the following
8. Wear protective clothing
Elastic protective clothing such as aprons, vests, shirts, skirts, thyroid protectors, and gloves, glasses, are used in stationary or moving environments without shielding
they will never completely stop X-rays, only reduce them to an acceptable level
At 100 kV, 3.2% of the beam will pass through a 0.5 mm lead shirt, and at 70 kV 0.36% will pass.
if the physician wears a lead apron and gloves but stands in the way of the primary radiation path, full security is not guaranteed
A lead apron used correctly can offer 80% protection to active blood -producing organs
A well -chosen apron should start just below the manubrium sternum to include the pubis symphysis index down to just above the knee.

## Slide 38: Every fluoroscopy operator should do the following
8. Wear protective clothing
g. for the protection to be effective, the face of the person must be turned toward the scattering source
h. The physician who is implementing the fluoroscopy procedure
should not turn their back on the beam
i. reinforced or equivalent materials that are filled with lead, copper, barium, and tungsten
j. These aprons should not be folded or wrinkled, and should not be thrown around randomly
k. addition, they should be checked every year, including vital areas to ensure there are no cracks or breaks
l. To protect against scattered radiation from the patient, two hanging
lead curtains next to the sides of the fluoroscopy tables can be used